## Results
From: | Laura Liermann, 312 Hosler Building
To: | Beth Herndon
Date | 2023-01-23
Basis | As-Received
Instrument | Thermo iCAP 7400, Inductively Coupled Plasma Emission Spectrometry (ICP-AES)
Method | Solution
Subject | Spectrochemical Characterization
Units | μg/ml |  | ppm = μg/ml = µg g-1 = µg/gm = mg/L
Det. Limit |  | 0.01 | 0.01 | 0.005 | 0.01 | 0.01 | 0.005 | 0.01 | 0.01
Our # | Your # | Al (ug/mL) | Ca (ug/mL) | Fe (ug/mL) | K (ug/mL) | Mg (ug/mL) | Mn (ug/mL) | Na (ug/mL) | P (ug/mL)
23-0156 | PhIr2022_156A | 0.0274 | 41.242 | 6.8014 | 0.0692 | 3.7712 | 1.5352 | 0.5693 | <0.01
23-0157 | PhIr2022_157A | 0.5304 | 5.3424 | 0.6131 | 0.3534 | 0.8555 | 0.1715 | 0.1545 | 0.0074
23-0158 | PhIr2022_158A | 0.5293 | 3.6483 | 0.5114 | 0.4965 | 0.7889 | 0.1173 | 0.2138 | 0.008
23-0159 | PhIr2022_159A | 0.5597 | 4.6265 | 0.5605 | 0.11 | 0.8583 | 0.0386 | 0.1941 | 0.0078
23-0160 | PhIr2022_160A | 0.2418 | 2.4327 | 0.1253 | 0.2606 | 0.6258 | 0.0241 | 0.202 | 0.0084
23-0161 | PhIr2022_161A | 0.227 | 2.1362 | 0.1498 | 0.2096 | 0.5521 | 0.0457 | 0.1735 | <0.01
23-0162 | PhIr2022_162A | 0.271 | 2.5443 | 0.3002 | 0.1075 | 0.644 | 0.2397 | 0.1878 | <0.01
23-0163 | PhIr2022_163A | 0.3006 | 2.7026 | 0.3273 | 0.1647 | 0.665 | 0.2099 | 0.2651 | 0.0068
23-0164 | PhIr2022_164A | 0.4254 | 2.7455 | 0.4717 | 0.3669 | 0.6179 | 0.2564 | 0.3274 | 0.0064
23-0165 | PhIr2022_165A | 0.289 | 2.1643 | 0.2591 | 0.0921 | 0.549 | 0.1648 | 0.1888 | <0.01
23-0166 | PhIr2022_166A | 0.5537 | 2.201 | 5.9253 | 0.0812 | 0.5782 | 0.2837 | 0.1527 | 0.0119
23-0167 | PhIr2022_167A | 0.3943 | 1.5235 | 3.9516 | 0.0294 | 0.4177 | 0.1236 | 0.211 | 0.0076
23-0168 | PhIr2022_168A | 0.3072 | 3.6178 | 0.1677 | 0.2748 | 0.6567 | 0.0702 | 0.2877 | 0.0078
23-0169 | PhIr2022_169A | 0.5431 | 8.821 | 11.3619 | 0.3353 | 1.7568 | 1.198 | 0.2282 | 0.0152
23-0170 | PhIr2022_170A | 0.3958 | 3.7509 | 1.9175 | 0.1249 | 0.7167 | 0.97284 | 0.3123 | <0.01
23-0171 | PhIr2022_171A | 0.2586 | 4.0968 | 0.2364 | 0.5343 | 1.0043 | 1.322 | 0.2036 | 0.0063
23-0172 | PhIr2022_172A | 0.791 | 21.2432 | 66.1955 | 0.7532 | 4.2207 | 51.2313 | 0.338 | <0.01
23-0173 | PhIr2022_173A | 0.4824 | 4.3084 | 22.5055 | 0.1685 | 1.1496 | 1.3857 | 0.3816 | 0.0105
23-0174 | PhIr2022_174A | -0.0008 | 0.0846 | <0.005 | 0.0648 | 0.0085 | <0.005 | 0.0393 | <0.01
23-0175 | PhIr2022_175A_10x | 0.0579 | 0.6403 | 0.1257 | 0.0305 | 0.0961 | 0.0299 | 0.0388 | <0.01
23-0176 | PhIr2022_176A | 0.424 | 4.6485 | 3.3418 | 0.3735 | 0.7733 | 1.7665 | 0.3379 | 0.0098
23-0177 | PhIr2022_177A_NOSAMPLE
23-0178 | PhIr2022_178A_10x | 0.08 | 0.6879 | 0.5628 | 0.0467 | 0.1084 | 0.1041 | 0.0522 | <0.01
23-0179 | PhIr2022_179A_10x | 0.06 | 0.5207 | 0.4107 | 0.0457 | 0.0929 | 0.1913 | 0.0395 | <0.01
23-0180 | PhIr2022_180A_10x | 0.0574 | 0.385 | 0.1617 | 0.0287 | 0.078 | 0.0601 | 0.0239 | <0.01
23-0181 | PhIr2022_181A_10x | 0.0873 | 0.5796 | 0.6347 | 0.046 | 0.0781 | 0.0947 | 0.0483 | <0.01
23-0182 | PhIr2022_182A_10x | 0.0727 | 0.4667 | 0.758 | 0.0286 | 0.067 | 0.0678 | 0.0398 | <0.01
23-0183 | PhIr2022_183A_10x | 0.1014 | 0.8387 | 0.899 | 0.0834 | 0.1183 | 0.4161 | 0.0637 | <0.01
23-0184 | PhIr2022_184A_10x | 0.0914 | 1.352 | 3.09 | 0.0369 | 0.1944 | 0.3395 | 0.0365 | <0.01
23-0185 | PhIr2022_185A_10x | 0.0901 | 0.6638 | 0.6664 | 0.0694 | 0.0912 | 0.1647 | 0.0568 | <0.01
23-0186 | PhIr2022_186A_10x | 0.0925 | 2.0402 | 10.7056 | 0.121 | 0.3688 | 7.0822 | 0.0408 | <0.01
23-0187 | PhIr2022_187A_10x | 0.2275 | 0.945 | 0.272 | 0.0383 | 0.1169 | 0.0915 | 0.0447 | <0.01
23-0188 | PhIr2022_188A_10x | 48.4864 | 20.6949 | 0.3514 | 0.0432 | 0.5596 | 0.0227 | 0.9514 | <0.01
23-0189 | PhIr2022_189A_10x | 0.0131 | 1.6889 | 0.0745 | 0.0168 | 0.2237 | 0.0196 | 0.0385 | <0.01
23-0190 | PhIr2022_190A_10x | 40.6884 | 8.1734 | 0.6524 | 0.0166 | 0.3165 | 0.0581 | 0.6758 | <0.01
23-0191 | PhIr2022_191A_10x | 13.8828 | 36.2827 | 0.3003 | 0.029 | 0.6583 | 0.0265 | 0.7729 | 0.0063
23-0192 | PhIr2022_192A_10x | 0.0209 | 3.5172 | 0.5214 | 0.0101 | 0.3557 | 0.1023 | 0.0533 | <0.01
23-0193 | Blank 1 | <0.01 | <0.01 | <0.005 | <0.01 | <0.01 | <0.005 | <0.01 | <0.01
23-0194 | Blank 2 | <0.01 | <0.01 | <0.005 | <0.01 | <0.01 | <0.005 | <0.01 | <0.01
23-0195 | Blank 3 | <0.01 | <0.01 | <0.005 | <0.01 | <0.01 | <0.005 | <0.01 | <0.01
EPA 200.7 QC standard |  | 20 | 20 | 20 | 100 | 20 | 20 | 20 | 100
EPA 200.7 diluted 1/100 |  | 0.2089 | 0.19814 | 0.1972 | 1.0219 | 0.20036 | 0.20068 | 0.2066 | 1.0086
EPA 200.7 diluted 1/20 |  | 1.0469 | 0.99694 | 0.99661 | 5.0776 | 1.0059 | 1.0121 | 1.0528 | 5.0392
EPA 200.7 diluted 1/10 |  | 2.104 | 1.9873 | 1.9801 | 10.138 | 1.9755 | 1.9953 | 1.9873 | 10.0159
EPA 200.7 diluted 1/4 |  | 5.3015 | 5.0392 | 4.9969 | 24.8075 | 4.9587 | 4.9541 | 5.0264 | 24.847
Note:
Results in ug/mL on as-received basis.
Samples were run on a Thermo iCAP 7400 ICP (inductively
coupled plasma emission spectrometer). Synthethic standards
from High Purity Standards were used to calibrate the results.
Note on C14: As-Received Basis = Values presented are the values in the sample as received by the lab. Sometimes samples are diluted to meet the parameters of the instrument however the results will always be reported back on an As-Received basis.

## Instruments
Instruments and Methods
Thermo iCAP 7400
Inductively Coupled Plasma Atomic Emission Spectrometry (ICP-AES)
Principle - solution is nebulized into a fine aerosol which is then introduced into a high energy plasma (temperature
between 7,000 TO 10,000k) The plasma excites the atoms and ions to higher energy levels. As they drop back down to
their normal levels, they do so by emitting energy wavelengths that are specific to the particular element. The wavelengths
of emitted energy are separated in an optical chamber. By monitoring what wavelengths are emitted, we can determine
what types of elements are present (qualitative determination). By comparing the intensity of the emitted wavelengths to
that of known standards, we can determine how much of each element is present (quantitative determination).
Technique Advantages
Quantitative determination of cations in a minimum of 5 ml of solution.
Levels of detection vary for each cation (generally from .002 to .01 ug/mL).
Some anions such as sulfur can also be determined.
Preferred technique for refractory cations such as Al, Si, W, Zr.
Capable of determining multiple elements simultaneously.
Other cations determined sequentially.
Data generated may be downloaded in Excel compatible format.
Typical Applications
Elemental studies in stream waters and acid mine drainage.
Cation measurements in fusion and digestion solutions.
Analyses of refractory elements in solution.
Other information
Recommended Reading
Handbook of Inductively Coupled Plasma Spectrometry 2nd edition, M. Thompson and J.N. Walsh, Blackie, 1989

## Definition of Terms
Definition of Terms
Basis | As-Recd = Values presented are the values in the sample as received by the lab.
Dry - Values presented are the values in the sample AFTER drying.
ASH = ash basis. In this case, the solid sample is ashed to eliminate volatiles and organics and the values presented are reported on the "ash" remaining.
Method | Technique(s) used to obtain the values presented
| LiBO2 | Original solid sample was dissolved via lithium metaborate fusion technique. Rock standards used to draw calibration curves.
| Acid Dig | Original solid sample was dissolved by acid digestion.
| Solution | Original solution samples were run directly. Solution standards were prepared from commercially certified stock solutions.
| Dilution | Original solution samples were diluted prior to running. Reported value is on AS-RECEIVED (original) basis.
Det. Limit | Approximate detection limit for the values presnted
Units | Units of measurement of the values presented. Units are typically in weight percent or parts per million (by weight)
Units | weight % - weight percent
| ppm - parts per million (by weight) | ppm = μg/ml = µg g-1 = µg/gm = mg/L
Other Terminology
Accuracy | ACCURACY is the comparison of the observed results to the "true value". True Value is
usually defined as the universally accepted value for a particular solid or solution
Precision | precision is the degree to which replicates agree with each other
(relative standard deviation)
Majors | "major" elements are those that make up significant parts of the sample. Typically, "majors" refer to concentrations in excess of 3 to 5 weight percent.
Minors | "minor" elements are those present in quantities ranging from 0.2 to 2 or 3 weight percent.
Traces | "trace" elements are those present in quantities less than approximately 0.10 weight percent.
Ultra-trace | "ultra-trace" elements are usually those present in quantities less than 100 parts per billion.
Concentration terminolgy
| ppm = parts per million = µg/mL = micrograms per milliliter = µg/gm = micrograms per gram |  | ppm = μg/ml = µg g-1 = µg/gm = mg/L
ppb = parts per billion = ng/mL = nanograms per milliliter = ng/gm = nanograms per gram
ppt = parts per trillion = pg/mL = picograms per milliliter = pg/gm = picograms per gram